

Best Practice – OMNIA Bridge

16.08.2017

─

Emanuele Magliozzi

Cosmed Asia Pacific

25 Dickson Ave

2064 Artarmon NSW

Contents

[Changes and Compatibility 2](#_Toc490826641)

[Overview 3](#_Toc490826642)

[Goals 3](#_Toc490826643)

[Pre-requisites 3](#_Toc490826644)

[Bridge setup 3](#_Toc490826645)

[Best Practice 3](#_Toc490826646)

[OMNIA 4](#_Toc490826647)

[Code Documentation 4](#_Toc490826648)

# Changes and Compatibility

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Date** | **Author** | **Notes** | **OCP v** | **BP API v** |
| 16/08/17 | Emanuele | First Draft | 0.4 | 1.0.6 |
|  |  |  |  |  |
|  |  |  |  |  |
|  |  |  |  |  |
|  |  |  |  |  |
|  |  |  |  |  |
|  |  |  |  |  |

# Overview

BP-OMNIA Bridge is a software (in C#) that allows Best Practice and OMNIA to communicate and to exchange personal details and spirometry results. The user can control and double check the data exchange using a basic graphic interface.

# Goals

1. Transfer person details from BP into OMNIA to create/select the patient/visit card
2. Transfer spirometry results plus PDF report into BP from OMNIA
3. Have this as simple and automatic as possible

# Pre-requisites

There is no control on OMNIA and BP installation and settings therefore the 2 software’s need to be installed and setup as suggested by the manufactures.

Furthermore, the software has been developed and tested using the following software versions:

* OMNIA v1.6.3
* BP Premier “Lava Edition” (1.8.5.754) or higher

# Bridge setup

In order to have the bridge working BP and OMNIA need to be setup as follow:

### Best Practice

As per documentation provided by manufacturer, part of the BP register need to setup in order to have the path of the file \*.exe of the bridge. This will allow BP to run the bridge when Utility/Spirometry is clicked. This is automatically set up during the installation of the bridge software. Here how it should look like:

[HKEY\_LOCAL\_MACHINE\SOFTWARE\Best Practice Software\Best Practice\Device **OR**

[HKEY\_LOCAL\_MACHINE\SOFTWARE\Wow6432Node\Best Practice Software\Best Practice\Device]

"SpiroPath" = ”path of the bridge \*.exe file”

"SpiroClass" = class name of the Form (something like that WindowsForms10.Window.8.app.0.141b42a\_r10\_ad1 which could be find using Spy++)

"SpiroName"= BP - OMNIA Bridge

### OMNIA

The bridge needs to login with specific rights so a new user needs to be created. Below the procedure:

1. Login as ***Admin***
2. Under ***Utility/User Accounts***, click on ***Users Management***
3. Click ***Add User***
4. Fill the form with:
   1. Username -> ***ocp***
   2. First Name -> ***BP Bridge***
   3. Unchecked ***Change PWD @ 1st Login***
   4. Checked ***PWD never expires***
5. Select ***PFT Technicians*** as role
6. Logout
7. Type ***ocp*** as username and click Login
8. Insert as password and confirm password ***bp*** and click OK

**+ flag in file ACS.XML as per OCP docs page 10**

# Code Documentation

The project is divided in 3 main parts:

1. Form
2. Omnia Command
   1. Command List
   2. Command Utility
3. Utility